## List of Mixes and Stock Sol
List of mixes and stock solutions (-80°C)
[5 empty] | LowIS Solutions |  |  |  | HighIS Solutions
Mix | Name | Storage Number | Concentration | Solvent | LowIS Volume (uL) | Formic Acid Volume (For 1% solution)(uL) | Final volume (uL) | Volume of ACN(uL) | HighIS Volume (uL) | Formic Acid Volume (For 1% solution)(uL) | Final volume (uL) | Volume of ACN (uL)
1 | 3-Phosphoglyceric Acid | 8 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| Acetylhistidine | 256 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Betaine | 152 | 1 mg/mL | H2O | 20 |  |  |  | 200
| C3 Carnitine | 79 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| CDP | 246 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Creatine | 42 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Dimethylglycine | 20 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glycerol | 171 | 1 mg/mL | H2O (1,6µL in 2000µL) | 20 |  |  |  | 200
| L-Glutamic Acid | 174 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Myo-Inositol | 193 | 1 mg/mL | H2O | 20 |  |  |  | 200
| propionic acid | 278 | 1mg/mL | H2O | 20 |  |  |  | 200
| Pyruvic Acid | 263 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Suberic Acid | 220 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Uric acid | 204 | 1 mg/mL | 2mM NaOH:H20 (2:1) + 5µL 0,3M NaOH | 20 |  |  |  | 200
| Xanthine | 205 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 |  |  |  | 200
2 | 1-Methylxanthine | 150 | 1 mg/mL | 1mM NaOH | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| 3-Nitrotyrosine | 250 | 1 mg/mL | DMSO | 20 |  |  |  | 200
| Ascorbic Acid | 153 | 1 mg/mL | H2O | 20 |  |  |  | 200
| C4 Carnitine | 80 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| cGMP | 237 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Fructose | 30 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glycine | 168 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Histidine | 3 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Hydroxyproline | 41 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Hypoxanthine | 58 | 1 mg/mL | H2O/Formic acid: H2O (2:1) | 20 |  |  |  | 200
| Lactic acid | 186 | 1 mg/mL | H2O | 20 |  |  |  | 200
| L-Kynurenine | 55 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Phosphocreatine | 31 | 1 mg/mL | H2O | 20 |  |  |  | 200
| UDP-glucuronate | 95 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Valine | 219 | 1 mg/mL | H2O | 20 |  |  |  | 200
3 | 5-Oxoproline | 19 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| Acetyl-CoA | 126 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Adenylosuccinic Acid | 149 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Alanine | 178 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Citrulline | 15 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Cytosine | 44 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glucose | 167 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glutathione Oxidized | 253 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glycolic acid | 165 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Guanine | 39 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 |  |  |  | 200
| Isoleucine | 182 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Malonic acid | 26 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Niacinamide | 23 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Palmitoylcarnitine | 82 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| UDP | 206 | 1 mg/mL | H2O | 20 |  |  |  | 200
4 | Acetylalanine | 67 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| Allantoin | 69 | 1 mg/mL | H2O | 20 |  |  |  | 200
| AMP | 154 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Gamma-Aminobutyric acid | 1 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glutamine | 108 | 1 mg/mL | H2O | 20 |  |  |  | 200
| GMP | 239 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Leucine | 214 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Phenylethylamine | 100 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Propylene glycol | 249 | 1 mg/mL | H2O (2µL in 2000µL) | 20 |  |  |  | 200
| Putrescine | 47 | 1 mg/mL | H2O (2,3µL in 2000µL) | 20 |  |  |  | 200
| Pyridoxine | 66 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Riboflavin | 14 | 1 mg/mL | DMSO | 20 |  |  |  | 200
| Ribose | 198 | 1 mg/mL | H2O | 20 |  |  |  | 200
| SAMe | 98 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Sedoheptulose-7-phosphate | 129 | 1 mg/mL | H2O | 20 |  |  |  | 200
5 | 1,3-Dimethyluric acid | 90 | 1 mg/mL | DMSO | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| 2-Ketobutyric acid | 52 | 1 mg/mL | H2O (+33°C) | 20 |  |  |  | 200
| 5-MTA | 102 | 1 mg/mL | H2O | 20 |  |  |  | 200
| CoA | 75 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Corticosterone | 247 | 1 mg/mL | MeOH | 20 |  |  |  | 200
| dADP | 244 | 1 mg/mL | H2O | 20 |  |  |  | 200
| dCDP | 233 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glycerophospho-inositol | 137 | 0,5mg/mL | EtOH | 20 |  |  |  | 200
| GTP | 236 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Mannose | 49 | 1 mg/mL | H2O | 20 |  |  |  | 200
| NADP | 135 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Nicotinic acid | 50 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Orotic Acid | 43 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 |  |  |  | 200
| PABA | 4 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| Threonine | 180 | 1 mg/mL | H2O | 20 |  |  |  | 200
6 | 3-Hydroxybutyric Acid | 57 | 1 mg/mL | H2O (1,8µL in 2000µL) | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| Acetylglutamic acid | 109 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Acetylglycine | 209 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Arginine | 283 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Asparagine | 181 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Choline | 12 | 1 mg/mL | H2O | 20 |  |  |  | 200
| CMP | 163 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Cystine | 54 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 |  |  |  | 200
| Histamine | 85 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Methioninesulfoxide | 122 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Paracetamol | 93 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| Sphingosine-1-phosphate | 138 | 1 mg/mL | 0,3M NaOH | 20 |  |  |  | 200
| Sucrose | 38 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Tryptophan | 223 | 1 mg/mL | EtOH + 5µL Sodium Hidroxide 5mMM | 20 |  |  |  | 200
| UDP-acetylglucosamine | 101 | 1 mg/mL | H2O | 20 |  |  |  | 200
7 | 1-Methyluric acid | 130 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| 3-Methylhistidine | 132 | 1 mg/mL | H2O | 20 |  |  |  | 200
| ADMA | 114 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Caffeine | 11 | 1 mg/mL | H2O | 20 |  |  |  | 200
| CDP-choline | 64 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Creatinine | 211 | 1 mg/mL | H2O | 20 |  |  |  | 200
| dAMP | 221 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glutaric acid | 213 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glycero-phosphocholine | 96 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Methionine | 215 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Phenylpyruvic acid | 92 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| Serine | 185 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Sphingosine | 140 | 1 mg/mL | MeOH | 20 |  |  |  | 200
| Taurine | 46 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Threonic Acid | 224 | 1 mg/mL | H2O | 20 |  |  |  | 200
8 | 5,6-Dihydro Thymine | 159 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| Aspirin | 18 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| Biotin | 29 | 1 mg/mL | DMSO | 20 |  |  |  | 200
| C5 Carnitine | 81 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| Cystathionine | 113 | 1 mg/mL | 1M NaOH | 20 |  |  |  | 200
| dGDP | 225 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glyceric Acid | 61 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Inosine | 189 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Mevalonic Acid | 124 | 1 mg/mL | H2O | 20 |  |  |  | 200
| NADPH | 115 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Norepinephrine | 87 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Ornithine | 22 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Succinic Acid | 197 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Tyrosine | 176 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| Uracil | 56 | 1 mg/mL | H2O | 20 |  |  |  | 200
9 | ADP-Ribose | 119 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| CDP-ethanolamine | 143 | 1 mg/mL | H2O | 20 |  |  |  | 200
| CTP | 160 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Diethyl malonate | 255 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Eplerone | 289 | 0,82mg/mL | MeOH | 20 |  |  |  | 200
| Erythrose 4-phosphate | 127 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glucose 1-Phosphate | 103 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Guanosine | 172 | 1 mg/mL | H2O | 20 |  |  |  | 200
| IMP | 51 | 1 mg/mL | H2O | 20 |  |  |  | 200
| L-Cysteine | 62 | 1 mg/mL | H2O | 20 |  |  |  | 200
| N-Acetylserine | 267 | 1 mg/mL | H2O | 20 |  |  |  | 200
| NADH | 120 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Pipecolic acid | 158 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Proline | 179 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Thymidine | 65 | 1 mg/mL | H2O | 20 |  |  |  | 200
10 | Adenosine | 155 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| Citric Acid | 162 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Cytidine | 161 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Fumaric Acid | 173 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| GDP | 231 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glucosamine | 27 | 1 mg/mL | H2O | 20 |  |  |  | 200
| L-Aspartic Acid | 282 | 1 mg/mL | H2O | 20 |  |  |  | 200
| N-Acetylornithine | 141 | 1 mg/mL | MeOH+H20 (1:1) | 20 |  |  |  | 200
| Phosphoenolpyruvic Acid | 72 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Phosphorylethanolamine | 7 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Ribulose 5-Phosphate | 133 | 1 mg/mL | H2O | 20 |  |  |  | 200
| SAH | 83 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Salicylic Acid | 9 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Spermidine | 235 | 1 mg/mL | H2O (2,2µL in 2000µL) | 20 |  |  |  | 200
| Thymine | 32 | 1 mg/mL | H2O | 20 |  |  |  | 200
11 | 2-Phosphoglyceric Acid | 111 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| 3-Methylxanthine | 123 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 |  |  |  | 200
| ADP | 238 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Alpha-ketoisovaleric acid | 144 | 1 mg/mL | H2O | 20 |  |  |  | 200
| dCTP | 226 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Dihydroorotic Acid | 105 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glutathione Reduced | 169 | 1 mg/mL | H2O | 20 |  |  |  | 200
| L,L-Cyclo(leucylprolyl) | 257 | 1 mg/mL | H2O | 20 |  |  |  | 200
| L-Carnitine | 281 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Lysine | 184 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Malic Acid | 175 | 1 mg/mL | H2O | 20 |  |  |  | 200
| p-Tyramine | 94 | 1 mg/mL | H2O | 20 |  |  |  | 200
| SDMA | 121 | 1 mg/mL | MeOH | 20 |  |  |  | 200
| Tetrahydrobiopterin | 89 | 1 mg/mL | H2O | 20 |  |  |  | 200
| UDP-Glucose | 202 | 1 mg/mL | H2O | 20 |  |  |  | 200
12 | Argininosuccinic acid | 59 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| Carnosine | 21 | 1 mg/mL | H2O | 20 |  |  |  | 200
| dCMP | 232 | 1 mg/mL | H2O | 20 |  |  |  | 200
| dGMP | 228 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Folic acid | 5 | 1 mg/mL | DMSO | 20 |  |  |  | 200
| galactitol | 164 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glucose 6-Phosphate | 240 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Guanidoacetic acid | 45 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Homocysteine | 217 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Ketoleucine | 63 | 1 mg/mL | H2O (1,9µL in 2000µL) | 20 |  |  |  | 200
| Methotrexate | 254 | 1 mg/mL | DMSO | 20 |  |  |  | 200
| N-Formyl-L-methionine | 227 | 1 mg/mL | MeOH | 20 |  |  |  | 200
| Sebacic acid | 218 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| Thiamine | 2 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Tryptamine | 48 | 1 mg/mL | EtOH | 20 |  |  |  | 200
13 | 1-Methylhistidine | 116 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| 8-Oxo-2-Deoxyguanosine | 252 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Acetylneuraminic Acid | 60 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Dopamine | 40 | 1 mg/mL | H2O | 20 |  |  |  | 200
| dUTP | 230 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Fructose-1,6-Diphosphate | 34 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Ibuprofen | 74 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| Indoleacetic acid | 13 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| N-Acetyl-beta-alanine | 259 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Pantothenic Acid | 6 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Phenylacetic acid | 10 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Phosphorylcholine | 25 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Sorbitol | 196 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Spermine | 234 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Uridine | 207 | 1 mg/mL | H2O | 20 |  |  |  | 200
14 | 1,5-anhydro D-glucitol | 251 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| 2-Ketoglutaric acid | 192 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Acetylcarnitine | 190 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| ADP-Glucose | 88 | 1 mg/mL | H2O | 20 |  |  |  | 200
| dGTP | 242 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glyceraldehyde 2-phosphate | 142 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glycerol-3-Phosphate | 170 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Indolelactic acid | 110 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Isocitric Acid | 188 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Levodopa | 84 | 1 mg/mL | H2O | 20 |  |  |  | 200
| NAD | 71 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Pyridoxic Acid | 97 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 |  |  |  | 200
| Serotonin | 70 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Ureidopropionic acid | 107 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Xylulose 5-Phosphate | 139 | 1 mg/mL | H2O | 20 |  |  |  | 200
15 | 2-Methylcitric acid | 117 | 1 mg/mL | H2O | 20 | 40 | 240 | 3760 | 200 | 40 | 2040 | 1960
| alpha-Aminoadipic acid | 106 | 1 mg/mL | H2O | 20 |  |  |  | 200
| FAD | 73 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glucosamine-6-Phosphate | 37 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Glucuronic Acid | 33 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Isovalerylglycine | 78 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| N-alpha-Acetyl-L-arginine | 86 | 1 mg/mL | DMSO | 20 |  |  |  | 200
| Phenylalanine | 177 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Phosphoserine | 16 | 1 mg/mL | H2O | 20 |  |  |  | 200
| p-Hydroxyphenylacetic acid | 35 | 1 mg/mL | H2O | 20 |  |  |  | 200
16 | 1,7-Dimethyluric acid | 146 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 | 40 | 180 | 3820 | 200 | 40 | 1440 | 2560
| 6-Phosphogluconic Acid | 77 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Acetyl-Glucosamine | 17 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Acetylmethionine | 91 | 1 mg/mL | MeOH | 20 |  |  |  | 200
| Dihydroxyacetone phosphate | 243 | 1 mg/mL | H2O | 20 |  |  |  | 200
| homovanillic acid | 166 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Quinolinic acid | 24 | 1 mg/mL | H2O | 20 |  |  |  | 200
17 | Succinylacetone | 134 | 1 mg/mL | H2O | 20 | 40 | 340 | 3660 | 200 | 40 | 3040 | 960
| dATP | 245 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Nicotine | 131 | 1 mg/mL | EtOH | 20 |  |  |  | 200
| Indoxyl sulfate | 53 | 1 mg/mL | H2O | 20 |  |  |  | 200
| 7-Methylxanthine | 147 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 |  |  |  | 200
| Acetylaspartic acid | 128 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Epinephrine | 36 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Xanthosine | 118 | 1 mg/mL | H2O + 5µL Sodium Hidroxide 5mM | 20 |  |  |  | 200
| 3-Hydroxy-DL-kynurenine | 216 | 1 mg/mL | DMSO | 20 |  |  |  | 200
| Gluconic Acid | 212 | 1 mg/mL | H2O | 20 |  |  |  | 200
| alpha-Lactose | 258 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Phosphoribosyl pyrophosphate | 136 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Ribose 5-Phosphate | 248 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Oxalacetic acid | 194 | 1 mg/mL | H2O | 20 |  |  |  | 200
| UTP | 201 | 1 mg/mL | H2O | 20 |  |  |  | 200
18 | 2,3-Diphosphoglyceric acid | 104 | NOT ORDERED | NOT ORDERED | 0 | 40 | 120 | 3880 | 200 | 40 | 1040 | 2960
| 2-Hydroxybutyric acid | 112 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Fructose 6-Phosphate | 28 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Octanoic-d15 acid | 76 | 1 mg/mL | DMSO | 20 |  |  |  | 200
| Succinyl-CoA | 145 | 1 mg/mL | H2O | 20 |  |  |  | 200
19 | Adenine | 261 | 1 mg/mL | H2O/Formic acid: H2O (2:1)
| ATP | 210 | 1 mg/mL | H2O
| cAMP | 156 | NOT ORDERED | NOT ORDERED | [6 empty] | LowIS Solutions |  |  |  | HighIS Solutions
| dUMP | 241 | NOT ORDERED | NOT ORDERED | [6 empty] | LowIS Volume (uL) | Formic Acid Volume (For 1% solution)(uL) | Final volume (uL) | Volume of CAN (uL) | HighIS Volume (uL) | Formic Acid Volume (For 1% solution)(uL) | Final volume (uL) | Volume of ACN(uL)
| Maleic acid | 183 | NOT ORDERED | NOT ORDERED |  | 19 | Adenine | 261 | 1 mg/mL | H2O/Formic acid: H2O (2:1) | 20 | 40 | 180 | 3820 | 200 | 40 | 1440 | 2560
| Methylmalonic acid | 191 | NOT ORDERED | NOT ORDERED |  |  | ATP | 210 | 1 mg/mL | H2O | 20 |  |  |  | 200
| UMP | 203 | 1 mg/mL | H2O |  |  | Methylamine | 265 | 1 mg/mL | H2O | 20 |  |  |  | 200
20 | 3-Chloro-5-(trifluoromethyl)-2-pyrindol | 269 | No solution prepared | No solution prepared |  |  | Trimethylamine | 222 | 1 mg/mL | H2O | 20 |  |  |  | 200
| 3,5-Dichloro-2-hydroxypyridine | 270 | No solution prepared | No solution prepared |  |  | Acetic Acid | 151 | 1 mg/mL | H2O | 20 |  |  |  | 200
| 3,5,6-Trichloro-2-pyrinidol | 273 | No solution prepared | No solution prepared |  |  | Homocysteinethiolactone | 260 | 1 mg/mL | DMSO | 20 |  |  |  | 200
| Methylamine | 265 | 1 mg/mL | H2O |  |  | UMP | 203 | 1 mg/mL | H2O | 20 |  |  |  | 200
| Trimethylamine | 222 | 1 mg/mL | H2O |  | 20 | 3-Chloro-5-(trifluoromethyl)-2-pyrindol | 269 | No solution prepared | No solution prepared | 20 | 40 | 140 | 3860 | 200 | 40 | 1040 | 2960
| Acetic Acid | 151 | 1 mg/mL | H2O |  |  | 3,5-Dichloro-2-hydroxypyridine | 270 | No solution prepared | No solution prepared | 20 |  |  |  | 200
| Homocysteinethiolactone | 260 | 1 mg/mL | DMSO |  |  | 3,5,6-Trichloro-2-pyrinidol | 273 | No solution prepared | No solution prepared | 20 |  |  |  | 200
| 2,6-Di-tert-butyl-4-methylphenol | 254 | No solution prepared | No solution prepared |  |  | 2,6-Di-tert-butyl-4-methylphenol | 254 | No solution prepared | No solution prepared | 20 |  |  |  | 200
| 7-Aminoheptanoic acid | 292 | No solution prepared | No solution prepared |  |  | 7-Aminoheptanoic acid | 292 | No solution prepared | No solution prepared | 20 |  |  |  | 200

## Calculations for Mix20
|  | Molecular Weight | For 1mL of 10mg/mL] | For 1mL of 1mg/mL]
3-Chloro-5-(trifluoromethyl)-2-pyrindol | 269 | 197.54
3,5-Dichloro-2-hydroxypyridine | 270 | 163.99
3,5,6-Trichloro-2-pyrinidol | 273 | 198.43
2,6-Di-tert-butyl-4-methylphenol | 254 | 220.35
7-Aminoheptanoic acid | 292 | 145.2

## Experiment Description
Experiment Descripition
This experiment aims to provide insights about RT variations due to matrix effects and/or concentration level of around 200 compounds by comparing the observed RT in water/ACN and in serum in 2 different concentrations. Moreover, fragmentantion patterns will be obtained (MS/MS) in both matrixes using data dependent acquisition (or IDA) in 2 different colision energies (20V and 30V), only for the most concentrated sample
Sample types
For each mix previously described, four types of samples will be prepared:
Batch 1: Water/ACN | LowIS | Mix of standards prepared in Water/ACN (50:50) at 5ppm
| HighIS | Mix of standards prepared in Water/ACN (50:50) at 5ppm
Batch 2: Serum | QC_LowIS | QC sample prepared exactly according to untargeted protocol
| QC_HighIS | QC sample prepared according to untargeted protocol using HighIS solution
Batch 1: Water/ACN | Blank#_Water_ACN | Mixture of Water/ACN (50:50) used to dilute the standards in LowIS and HighIS samples
Batch 2: Serum | Blank#_QC | A serum sample prepared according to untargeted protocol but with no IS added
# represents a number from 01 to 03
Experiment Info
1. Data acquisition will be divided in two different batches according to the different matrixes since column conditioning should be different
2. Each type of sample will be analysed in triplicate (triplicates of sample preparation, not analytical)
3. DDA data will be acquired for HighIS and QC_HighIS samples only. From the triplicate,sample with _2 will be selected and analysed under 20V and 30V
4. For serum batch, Blank_Water_ACN will be injected as a blank and Blank_QC will be used for column conditioning at the beggining. For Water/ACN batch, only Blank_Water_ACN will be injected
5. Samples will be injected randomly (see sheet MS Analysis Plan)
Samples List
| Batch Water/ACN |  | Batch Serum |  | Total per batch: | Divided in: | Samples per plate:
[5 empty] | 120 samples | 2 plates | 63 (60 samples + 3 Blanks)
| LowIS_Mix01_1 | HighIS_Mix01_1 | QC_LowIS_Mix01_1 | QC_HighIS_Mix01_1
| LowIS_Mix01_2 | HighIS_Mix01_2 | QC_LowIS_Mix01_2 | QC_HighIS_Mix01_2
| LowIS_Mix01_3 | HighIS_Mix01_3 | QC_LowIS_Mix01_3 | QC_HighIS_Mix01_3
| LowIS_Mix02_1 | HighIS_Mix02_1 | QC_LowIS_Mix02_1 | QC_HighIS_Mix02_1
| LowIS_Mix02_2 | HighIS_Mix02_2 | QC_LowIS_Mix02_2 | QC_HighIS_Mix02_2
| LowIS_Mix02_3 | HighIS_Mix02_3 | QC_LowIS_Mix02_3 | QC_HighIS_Mix02_3
| LowIS_Mix03_1 | HighIS_Mix03_1 | QC_LowIS_Mix03_1 | QC_HighIS_Mix03_1
| LowIS_Mix03_2 | HighIS_Mix03_2 | QC_LowIS_Mix03_2 | QC_HighIS_Mix03_2
| LowIS_Mix03_3 | HighIS_Mix03_3 | QC_LowIS_Mix03_3 | QC_HighIS_Mix03_3
| LowIS_Mix04_1 | HighIS_Mix04_1 | QC_LowIS_Mix04_1 | QC_HighIS_Mix04_1
| LowIS_Mix04_2 | HighIS_Mix04_2 | QC_LowIS_Mix04_2 | QC_HighIS_Mix04_2
| LowIS_Mix04_3 | HighIS_Mix04_3 | QC_LowIS_Mix04_3 | QC_HighIS_Mix04_3
| LowIS_Mix05_1 | HighIS_Mix05_1 | QC_LowIS_Mix05_1 | QC_HighIS_Mix05_1
| LowIS_Mix05_2 | HighIS_Mix05_2 | QC_LowIS_Mix05_2 | QC_HighIS_Mix05_2
| LowIS_Mix05_3 | HighIS_Mix05_3 | QC_LowIS_Mix05_3 | QC_HighIS_Mix05_3
| LowIS_Mix06_1 | HighIS_Mix06_1 | QC_LowIS_Mix06_1 | QC_HighIS_Mix06_1
| LowIS_Mix06_2 | HighIS_Mix06_2 | QC_LowIS_Mix06_2 | QC_HighIS_Mix06_2
| LowIS_Mix06_3 | HighIS_Mix06_3 | QC_LowIS_Mix06_3 | QC_HighIS_Mix06_3
| LowIS_Mix07_1 | HighIS_Mix07_1 | QC_LowIS_Mix07_1 | QC_HighIS_Mix07_1
| LowIS_Mix07_2 | HighIS_Mix07_2 | QC_LowIS_Mix07_2 | QC_HighIS_Mix07_2
| LowIS_Mix07_3 | HighIS_Mix07_3 | QC_LowIS_Mix07_3 | QC_HighIS_Mix07_3
| LowIS_Mix08_1 | HighIS_Mix08_1 | QC_LowIS_Mix08_1 | QC_HighIS_Mix08_1
| LowIS_Mix08_2 | HighIS_Mix08_2 | QC_LowIS_Mix08_2 | QC_HighIS_Mix08_2
| LowIS_Mix08_3 | HighIS_Mix08_3 | QC_LowIS_Mix08_3 | QC_HighIS_Mix08_3
| LowIS_Mix09_1 | HighIS_Mix09_1 | QC_LowIS_Mix09_1 | QC_HighIS_Mix09_1
| LowIS_Mix09_2 | HighIS_Mix09_2 | QC_LowIS_Mix09_2 | QC_HighIS_Mix09_2
| LowIS_Mix09_3 | HighIS_Mix09_3 | QC_LowIS_Mix09_3 | QC_HighIS_Mix09_3
| LowIS_Mix10_1 | HighIS_Mix10_1 | QC_LowIS_Mix10_1 | QC_HighIS_Mix10_1
| LowIS_Mix10_2 | HighIS_Mix10_2 | QC_LowIS_Mix10_2 | QC_HighIS_Mix10_2
| LowIS_Mix10_3 | HighIS_Mix10_3 | QC_LowIS_Mix10_3 | QC_HighIS_Mix10_3
| LowIS_Mix11_1 | HighIS_Mix11_1 | QC_LowIS_Mix11_1 | QC_HighIS_Mix11_1
| LowIS_Mix11_2 | HighIS_Mix11_2 | QC_LowIS_Mix11_2 | QC_HighIS_Mix11_2
| LowIS_Mix11_3 | HighIS_Mix11_3 | QC_LowIS_Mix11_3 | QC_HighIS_Mix11_3
| LowIS_Mix12_1 | HighIS_Mix12_1 | QC_LowIS_Mix12_1 | QC_HighIS_Mix12_1
| LowIS_Mix12_2 | HighIS_Mix12_2 | QC_LowIS_Mix12_2 | QC_HighIS_Mix12_2
| LowIS_Mix12_3 | HighIS_Mix12_3 | QC_LowIS_Mix12_3 | QC_HighIS_Mix12_3
| LowIS_Mix13_1 | HighIS_Mix13_1 | QC_LowIS_Mix13_1 | QC_HighIS_Mix13_1
| LowIS_Mix13_2 | HighIS_Mix13_2 | QC_LowIS_Mix13_2 | QC_HighIS_Mix13_2
| LowIS_Mix13_3 | HighIS_Mix13_3 | QC_LowIS_Mix13_3 | QC_HighIS_Mix13_3
| LowIS_Mix14_1 | HighIS_Mix14_1 | QC_LowIS_Mix14_1 | QC_HighIS_Mix14_1
| LowIS_Mix14_2 | HighIS_Mix14_2 | QC_LowIS_Mix14_2 | QC_HighIS_Mix14_2
| LowIS_Mix14_3 | HighIS_Mix14_3 | QC_LowIS_Mix14_3 | QC_HighIS_Mix14_3
| LowIS_Mix15_1 | HighIS_Mix15_1 | QC_LowIS_Mix15_1 | QC_HighIS_Mix15_1
| LowIS_Mix15_2 | HighIS_Mix15_2 | QC_LowIS_Mix15_2 | QC_HighIS_Mix15_2
| LowIS_Mix15_3 | HighIS_Mix15_3 | QC_LowIS_Mix15_3 | QC_HighIS_Mix15_3
| LowIS_Mix16_1 | HighIS_Mix16_1 | QC_LowIS_Mix16_1 | QC_HighIS_Mix16_1
| LowIS_Mix16_2 | HighIS_Mix16_2 | QC_LowIS_Mix16_2 | QC_HighIS_Mix16_2
| LowIS_Mix16_3 | HighIS_Mix16_3 | QC_LowIS_Mix16_3 | QC_HighIS_Mix16_3
| LowIS_Mix17_1 | HighIS_Mix17_1 | QC_LowIS_Mix17_1 | QC_HighIS_Mix17_1
| LowIS_Mix17_2 | HighIS_Mix17_2 | QC_LowIS_Mix17_2 | QC_HighIS_Mix17_2
| LowIS_Mix17_3 | HighIS_Mix17_3 | QC_LowIS_Mix17_3 | QC_HighIS_Mix17_3
| LowIS_Mix18_1 | HighIS_Mix18_1 | QC_LowIS_Mix18_1 | QC_HighIS_Mix18_1
| LowIS_Mix18_2 | HighIS_Mix18_2 | QC_LowIS_Mix18_2 | QC_HighIS_Mix18_2
| LowIS_Mix18_3 | HighIS_Mix18_3 | QC_LowIS_Mix18_3 | QC_HighIS_Mix18_3
| LowIS_Mix19_1 | HighIS_Mix19_1 | QC_LowIS_Mix19_1 | QC_HighIS_Mix19_1
| LowIS_Mix19_2 | HighIS_Mix19_2 | QC_LowIS_Mix19_2 | QC_HighIS_Mix19_2
| LowIS_Mix19_3 | HighIS_Mix19_3 | QC_LowIS_Mix19_3 | QC_HighIS_Mix19_3
| LowIS_Mix20_1 | HighIS_Mix20_1 | QC_LowIS_Mix20_1 | QC_HighIS_Mix20_1
| LowIS_Mix20_2 | HighIS_Mix20_2 | QC_LowIS_Mix20_2 | QC_HighIS_Mix20_2
| LowIS_Mix20_3 | HighIS_Mix20_3 | QC_LowIS_Mix20_3 | QC_HighIS_Mix20_3

## MS Analysis Plan and Plate Map
MS Analysis Plan and Plate Map
Water/ACN Batch |  |  |  |  | Serum Batch | [5 empty] | * The final file name will be generated following this names + POS/NEG
* The final file name will be generated following this names + POS/NEG |  |  |  |  | * The final file name will be generated following this names + POS/NEG
Full List | Randomized | Final Worklist (with Blanks and MS/MS) | Plate Number | Deepwel Plate Position | Full List | Randomized | Final Worklist (with Blanks and MS/MS) |  | Plate Number | Deepwel Plate Position | Direction of well counting in the software
LowIS_Mix01_1 | LowIS_Mix06_3 | Blank1_Water_ACN_1 | 1 | 12 | QC_LowIS_Mix01_1 | QC_LowIS_Mix14_2 | Blank01_QC_1 | *Fifteen injections as in the equilibration mix | 1 | 12 | Plate Position Map
LowIS_Mix01_2 | HighIS_Mix13_3 | Blank2_Water_ACN_1 | 1 | 24 | QC_LowIS_Mix01_2 | QC_HighIS_Mix04_3 | Blank02_QC_1 |  | 1 | 24 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
LowIS_Mix01_3 | HighIS_Mix09_2 | Blank3_Water_ACN_1 | 1 | 36 | QC_LowIS_Mix01_3 | QC_LowIS_Mix10_1 | Blank03_QC_1 |  | 1 | 36 | A | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
LowIS_Mix02_1 | HighIS_Mix20_1 | Blank1_Water_ACN_2 | 1 | 12 | QC_LowIS_Mix02_1 | QC_LowIS_Mix15_3 | Blank01_QC_2 |  | 1 | 12 | B | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24
LowIS_Mix02_2 | LowIS_Mix13_3 | Blank2_Water_ACN_2 | 1 | 24 | QC_LowIS_Mix02_2 | QC_HighIS_Mix19_3 | Blank02_QC_2 |  | 1 | 24 | C | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36
LowIS_Mix02_3 | LowIS_Mix12_2 | Blank3_Water_ACN_2 | 1 | 36 | QC_LowIS_Mix02_3 | QC_LowIS_Mix13_1 | Blank03_QC_2 |  | 1 | 36 | D | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48
LowIS_Mix03_1 | LowIS_Mix14_3 | Blank1_Water_ACN_3 | 1 | 12 | QC_LowIS_Mix03_1 | QC_LowIS_Mix04_1 | Blank01_QC_3 |  | 1 | 12 | E | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60
LowIS_Mix03_2 | HighIS_Mix15_2 | Blank2_Water_ACN_3 | 1 | 24 | QC_LowIS_Mix03_2 | QC_HighIS_Mix13_3 | Blank02_QC_3 |  | 1 | 24 | F | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72
LowIS_Mix03_3 | LowIS_Mix02_2 | Blank3_Water_ACN_3 | 1 | 36 | QC_LowIS_Mix03_3 | QC_LowIS_Mix01_2 | Blank03_QC_3 |  | 1 | 36 | G | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84
LowIS_Mix04_1 | LowIS_Mix19_2 | LowIS_Mix06_3 | 1 | 30 | QC_LowIS_Mix04_1 | QC_LowIS_Mix18_3 | Blank01_QC_4 |  | 1 | 12 | H | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96
LowIS_Mix04_2 | LowIS_Mix10_3 | HighIS_Mix13_3 | 2 | 87 | QC_LowIS_Mix04_2 | QC_LowIS_Mix03_1 | Blank02_QC_4 |  | 1 | 24
LowIS_Mix04_3 | HighIS_Mix17_3 | HighIS_Mix09_2 | 1 | 81 | QC_LowIS_Mix04_3 | QC_LowIS_Mix14_1 | Blank03_QC_4 |  | 1 | 36 | Water/ACN Batch - Plate #1
LowIS_Mix05_1 | HighIS_Mix12_3 | HighIS_Mix20_1 | 2 | 70 | QC_LowIS_Mix05_1 | QC_LowIS_Mix08_3 | Blank01_QC_5 |  | 1 | 12 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
LowIS_Mix05_2 | LowIS_Mix09_1 | LowIS_Mix13_3 | 2 | 27 | QC_LowIS_Mix05_2 | QC_LowIS_Mix13_3 | Blank02_QC_5 |  | 1 | 24 | A | LowIS_Mix01_1 | LowIS_Mix02_1 | LowIS_Mix03_1 | LowIS_Mix04_1 | LowIS_Mix05_1 | LowIS_Mix06_1 | LowIS_Mix07_1 | LowIS_Mix08_1 | LowIS_Mix09_1 | LowIS_Mix10_1 |  | Blank01_Water_ACN
LowIS_Mix05_3 | LowIS_Mix12_3 | LowIS_Mix12_2 | 2 | 14 | QC_LowIS_Mix05_3 | QC_HighIS_Mix14_3 | Blank03_QC_5 |  | 1 | 36 | B | LowIS_Mix01_2 | LowIS_Mix02_2 | LowIS_Mix03_2 | LowIS_Mix04_2 | LowIS_Mix05_2 | LowIS_Mix06_2 | LowIS_Mix07_2 | LowIS_Mix08_2 | LowIS_Mix09_2 | LowIS_Mix10_2 |  | Blank02_Water_ACN
LowIS_Mix06_1 | HighIS_Mix03_1 | LowIS_Mix14_3 | 2 | 28 | QC_LowIS_Mix06_1 | QC_HighIS_Mix08_3 | QC_LowIS_Mix14_2 |  | 2 | 16 | C | LowIS_Mix01_3 | LowIS_Mix02_3 | LowIS_Mix03_3 | LowIS_Mix04_3 | LowIS_Mix05_3 | LowIS_Mix06_3 | LowIS_Mix07_3 | LowIS_Mix08_3 | LowIS_Mix09_3 | LowIS_Mix10_3 |  | Blank03_Water_ACN
LowIS_Mix06_2 | LowIS_Mix20_2 | HighIS_Mix15_2 | 2 | 77 | QC_LowIS_Mix06_2 | QC_HighIS_Mix10_3 | QC_HighIS_Mix04_3 |  | 1 | 88 | D
LowIS_Mix06_3 | LowIS_Mix07_1 | LowIS_Mix02_2 | 1 | 14 | QC_LowIS_Mix06_3 | QC_LowIS_Mix07_1 | QC_LowIS_Mix10_1 |  | 1 | 10 | E
LowIS_Mix07_1 | LowIS_Mix13_2 | LowIS_Mix19_2 | 2 | 21 | QC_LowIS_Mix07_1 | QC_LowIS_Mix11_1 | QC_LowIS_Mix15_3 |  | 2 | 29 | F | HighIS_Mix01_1 | HighIS_Mix02_1 | HighIS_Mix03_1 | HighIS_Mix04_1 | HighIS_Mix05_1 | HighIS_Mix06_1 | HighIS_Mix07_1 | HighIS_Mix08_1 | HighIS_Mix09_1 | HighIS_Mix10_1
LowIS_Mix07_2 | HighIS_Mix02_3 | LowIS_Mix10_3 | 1 | 34 | QC_LowIS_Mix07_2 | QC_HighIS_Mix04_1 | QC_HighIS_Mix19_3 |  | 2 | 93 | G | HighIS_Mix01_2 | HighIS_Mix02_2 | HighIS_Mix03_2 | HighIS_Mix04_2 | HighIS_Mix05_2 | HighIS_Mix06_2 | HighIS_Mix07_2 | HighIS_Mix08_2 | HighIS_Mix09_2 | HighIS_Mix10_2
LowIS_Mix07_3 | LowIS_Mix16_3 | HighIS_Mix17_3 | 2 | 91 | QC_LowIS_Mix07_3 | QC_HighIS_Mix04_2 | QC_LowIS_Mix13_1 |  | 2 | 3 | H | HighIS_Mix01_3 | HighIS_Mix02_3 | HighIS_Mix03_3 | HighIS_Mix04_3 | HighIS_Mix05_3 | HighIS_Mix06_3 | HighIS_Mix07_3 | HighIS_Mix08_3 | HighIS_Mix09_3 | HighIS_Mix10_3
LowIS_Mix08_1 | LowIS_Mix07_3 | HighIS_Mix12_3 | 2 | 86 | QC_LowIS_Mix08_1 | QC_HighIS_Mix11_2 | QC_LowIS_Mix04_1 |  | 1 | 4
LowIS_Mix08_2 | HighIS_Mix14_1 | LowIS_Mix09_1 | 1 | 9 | QC_LowIS_Mix08_2 | QC_HighIS_Mix10_2 | QC_HighIS_Mix13_3 |  | 2 | 87
LowIS_Mix08_3 | HighIS_Mix03_3 | LowIS_Mix12_3 | 2 | 26 | QC_LowIS_Mix08_3 | QC_LowIS_Mix04_2 | QC_LowIS_Mix01_2 |  | 1 | 13 | Water/ACN Batch - Plate #2
LowIS_Mix09_1 | HighIS_Mix11_2 | HighIS_Mix03_1 | 1 | 63 | QC_LowIS_Mix09_1 | QC_LowIS_Mix02_2 | QC_LowIS_Mix18_3 |  | 2 | 32 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
LowIS_Mix09_2 | HighIS_Mix19_3 | LowIS_Mix20_2 | 2 | 22 | QC_LowIS_Mix09_2 | QC_LowIS_Mix18_1 | QC_LowIS_Mix03_1 |  | 1 | 3 | A | LowIS_Mix11_1 | LowIS_Mix12_1 | LowIS_Mix13_1 | LowIS_Mix14_1 | LowIS_Mix15_1 | LowIS_Mix16_1 | LowIS_Mix17_1 | LowIS_Mix18_1 | LowIS_Mix19_1 | LowIS_Mix20_1
LowIS_Mix09_3 | HighIS_Mix06_2 | LowIS_Mix07_1 | 1 | 7 | QC_LowIS_Mix09_3 | QC_HighIS_Mix08_1 | QC_LowIS_Mix14_1 |  | 2 | 4 | B | LowIS_Mix11_2 | LowIS_Mix12_2 | LowIS_Mix13_2 | LowIS_Mix14_2 | LowIS_Mix15_2 | LowIS_Mix16_2 | LowIS_Mix17_2 | LowIS_Mix18_2 | LowIS_Mix19_2 | LowIS_Mix20_2
LowIS_Mix10_1 | HighIS_Mix18_3 | LowIS_Mix13_2 | 2 | 15 | QC_LowIS_Mix10_1 | QC_HighIS_Mix13_2 | QC_LowIS_Mix08_3 |  | 1 | 32 | C | LowIS_Mix11_3 | LowIS_Mix12_3 | LowIS_Mix13_3 | LowIS_Mix14_3 | LowIS_Mix15_3 | LowIS_Mix16_3 | LowIS_Mix17_3 | LowIS_Mix18_3 | LowIS_Mix19_3 | LowIS_Mix20_3
LowIS_Mix10_2 | LowIS_Mix07_2 | HighIS_Mix02_3 | 1 | 86 | QC_LowIS_Mix10_2 | QC_HighIS_Mix20_2 | QC_LowIS_Mix13_3 |  | 2 | 27 | D
LowIS_Mix10_3 | LowIS_Mix19_3 | LowIS_Mix16_3 | 2 | 30 | QC_LowIS_Mix10_3 | QC_LowIS_Mix08_2 | QC_HighIS_Mix14_3 |  | 2 | 88 | E
LowIS_Mix11_1 | LowIS_Mix02_3 | LowIS_Mix07_3 | 1 | 31 | QC_LowIS_Mix11_1 | QC_LowIS_Mix09_2 | QC_HighIS_Mix08_3 |  | 1 | 92 | F | HighIS_Mix11_1 | HighIS_Mix12_1 | HighIS_Mix13_1 | HighIS_Mix14_1 | HighIS_Mix15_1 | HighIS_Mix16_1 | HighIS_Mix17_1 | HighIS_Mix18_1 | HighIS_Mix19_1 | HighIS_Mix20_1
LowIS_Mix11_2 | LowIS_Mix08_2 | HighIS_Mix14_1 | 2 | 64 | QC_LowIS_Mix11_2 | QC_LowIS_Mix17_2 | QC_HighIS_Mix10_3 |  | 1 | 94 | G | HighIS_Mix11_2 | HighIS_Mix12_2 | HighIS_Mix13_2 | HighIS_Mix14_2 | HighIS_Mix15_2 | HighIS_Mix16_2 | HighIS_Mix17_2 | HighIS_Mix18_2 | HighIS_Mix19_2 | HighIS_Mix20_2
LowIS_Mix11_3 | LowIS_Mix04_1 | HighIS_Mix03_3 | 1 | 87 | QC_LowIS_Mix11_3 | QC_LowIS_Mix08_1 | QC_LowIS_Mix07_1 |  | 1 | 7 | H | HighIS_Mix11_3 | HighIS_Mix12_3 | HighIS_Mix13_3 | HighIS_Mix14_3 | HighIS_Mix15_3 | HighIS_Mix16_3 | HighIS_Mix17_3 | HighIS_Mix18_3 | HighIS_Mix19_3 | HighIS_Mix20_3
LowIS_Mix12_1 | HighIS_Mix08_1 | HighIS_Mix11_2 | 2 | 73 | QC_LowIS_Mix12_1 | QC_HighIS_Mix09_1 | QC_LowIS_Mix11_1 |  | 2 | 1
LowIS_Mix12_2 | HighIS_Mix12_2 | HighIS_Mix19_3 | 2 | 93 | QC_LowIS_Mix12_2 | QC_HighIS_Mix17_2 | QC_HighIS_Mix04_1 |  | 1 | 64
LowIS_Mix12_3 | LowIS_Mix14_2 | HighIS_Mix06_2 | 1 | 78 | QC_LowIS_Mix12_3 | QC_HighIS_Mix17_3 | QC_HighIS_Mix04_2 |  | 1 | 76 | Serum Batch - Plate #1
LowIS_Mix13_1 | HighIS_Mix10_2 | HighIS_Mix18_3 | 2 | 92 | QC_LowIS_Mix13_1 | QC_HighIS_Mix03_1 | QC_HighIS_Mix11_2 |  | 2 | 73 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
LowIS_Mix13_2 | HighIS_Mix01_3 | LowIS_Mix07_2 | 1 | 19 | QC_LowIS_Mix13_2 | QC_LowIS_Mix16_3 | QC_HighIS_Mix10_2 |  | 1 | 82 | A | QC_LowIS_Mix01_1 | QC_LowIS_Mix02_1 | QC_LowIS_Mix03_1 | QC_LowIS_Mix04_1 | QC_LowIS_Mix05_1 | QC_LowIS_Mix06_1 | QC_LowIS_Mix07_1 | QC_LowIS_Mix08_1 | QC_LowIS_Mix09_1 | QC_LowIS_Mix10_1 |  | Blank01_QC
LowIS_Mix13_3 | LowIS_Mix11_2 | LowIS_Mix19_3 | 2 | 33 | QC_LowIS_Mix13_3 | QC_HighIS_Mix15_1 | QC_LowIS_Mix04_2 |  | 1 | 16 | B | QC_LowIS_Mix01_2 | QC_LowIS_Mix02_2 | QC_LowIS_Mix03_2 | QC_LowIS_Mix04_2 | QC_LowIS_Mix05_2 | QC_LowIS_Mix06_2 | QC_LowIS_Mix07_2 | QC_LowIS_Mix08_2 | QC_LowIS_Mix09_2 | QC_LowIS_Mix10_2 |  | Blank02_QC
LowIS_Mix14_1 | HighIS_Mix04_1 | LowIS_Mix02_3 | 1 | 26 | QC_LowIS_Mix14_1 | QC_HighIS_Mix16_3 | QC_LowIS_Mix02_2 |  | 1 | 14 | C | QC_LowIS_Mix01_3 | QC_LowIS_Mix02_3 | QC_LowIS_Mix03_3 | QC_LowIS_Mix04_3 | QC_LowIS_Mix05_3 | QC_LowIS_Mix06_3 | QC_LowIS_Mix07_3 | QC_LowIS_Mix08_3 | QC_LowIS_Mix09_3 | QC_LowIS_Mix10_3 |  | Blank03_QC
LowIS_Mix14_2 | HighIS_Mix09_1 | LowIS_Mix08_2 | 1 | 20 | QC_LowIS_Mix14_2 | QC_HighIS_Mix20_3 | QC_LowIS_Mix18_1 |  | 2 | 8 | D
LowIS_Mix14_3 | HighIS_Mix02_2 | LowIS_Mix04_1 | 1 | 4 | QC_LowIS_Mix14_3 | QC_LowIS_Mix09_3 | QC_HighIS_Mix08_1 |  | 1 | 68 | E
LowIS_Mix15_1 | HighIS_Mix02_1 | HighIS_Mix08_1 | 1 | 68 | QC_LowIS_Mix15_1 | QC_LowIS_Mix15_1 | QC_HighIS_Mix13_2 |  | 2 | 75 | F | QC_HighIS_Mix01_1 | QC_HighIS_Mix02_1 | QC_HighIS_Mix03_1 | QC_HighIS_Mix04_1 | QC_HighIS_Mix05_1 | QC_HighIS_Mix06_1 | QC_HighIS_Mix07_1 | QC_HighIS_Mix08_1 | QC_HighIS_Mix09_1 | QC_HighIS_Mix10_1
LowIS_Mix15_2 | LowIS_Mix06_1 | HighIS_Mix12_2 | 2 | 74 | QC_LowIS_Mix15_2 | QC_HighIS_Mix14_2 | QC_HighIS_Mix20_2 |  | 2 | 82 | G | QC_HighIS_Mix01_2 | QC_HighIS_Mix02_2 | QC_HighIS_Mix03_2 | QC_HighIS_Mix04_2 | QC_HighIS_Mix05_2 | QC_HighIS_Mix06_2 | QC_HighIS_Mix07_2 | QC_HighIS_Mix08_2 | QC_HighIS_Mix09_2 | QC_HighIS_Mix10_2
LowIS_Mix15_3 | LowIS_Mix05_3 | LowIS_Mix14_2 | 2 | 16 | QC_LowIS_Mix15_3 | QC_HighIS_Mix18_2 | QC_LowIS_Mix08_2 |  | 1 | 20 | H | QC_HighIS_Mix01_3 | QC_HighIS_Mix02_3 | QC_HighIS_Mix03_3 | QC_HighIS_Mix04_3 | QC_HighIS_Mix05_3 | QC_HighIS_Mix06_3 | QC_HighIS_Mix07_3 | QC_HighIS_Mix08_3 | QC_HighIS_Mix09_3 | QC_HighIS_Mix10_3
LowIS_Mix16_1 | LowIS_Mix01_2 | HighIS_Mix10_2 | 1 | 82 | QC_LowIS_Mix16_1 | QC_LowIS_Mix13_2 | QC_LowIS_Mix09_2 |  | 1 | 21
LowIS_Mix16_2 | HighIS_Mix09_3 | HighIS_Mix01_3 | 1 | 85 | QC_LowIS_Mix16_2 | QC_HighIS_Mix17_1 | QC_LowIS_Mix17_2 |  | 2 | 19
LowIS_Mix16_3 | LowIS_Mix17_1 | LowIS_Mix11_2 | 2 | 13 | QC_LowIS_Mix16_3 | QC_HighIS_Mix09_2 | QC_LowIS_Mix08_1 |  | 1 | 8 | Serum Batch - Plate #2
LowIS_Mix17_1 | HighIS_Mix16_1 | HighIS_Mix04_1 | 1 | 64 | QC_LowIS_Mix17_1 | QC_LowIS_Mix11_3 | QC_HighIS_Mix09_1 |  | 1 | 69 |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
LowIS_Mix17_2 | LowIS_Mix11_3 | HighIS_Mix09_1 | 1 | 69 | QC_LowIS_Mix17_2 | QC_LowIS_Mix03_2 | QC_HighIS_Mix17_2 |  | 2 | 79 | A | QC_LowIS_Mix11_1 | QC_LowIS_Mix12_1 | QC_LowIS_Mix13_1 | QC_LowIS_Mix14_1 | QC_LowIS_Mix15_1 | QC_LowIS_Mix16_1 | QC_LowIS_Mix17_1 | QC_LowIS_Mix18_1 | QC_LowIS_Mix19_1 | QC_LowIS_Mix20_1
LowIS_Mix17_3 | LowIS_Mix10_2 | HighIS_Mix02_2 | 1 | 74 | QC_LowIS_Mix17_3 | QC_LowIS_Mix16_2 | QC_HighIS_Mix17_3 |  | 2 | 91 | B | QC_LowIS_Mix11_2 | QC_LowIS_Mix12_2 | QC_LowIS_Mix13_2 | QC_LowIS_Mix14_2 | QC_LowIS_Mix15_2 | QC_LowIS_Mix16_2 | QC_LowIS_Mix17_2 | QC_LowIS_Mix18_2 | QC_LowIS_Mix19_2 | QC_LowIS_Mix20_2
LowIS_Mix18_1 | HighIS_Mix05_2 | HighIS_Mix02_1 | 1 | 62 | QC_LowIS_Mix18_1 | QC_HighIS_Mix12_3 | QC_HighIS_Mix03_1 |  | 1 | 63 | C | QC_LowIS_Mix11_3 | QC_LowIS_Mix12_3 | QC_LowIS_Mix13_3 | QC_LowIS_Mix14_3 | QC_LowIS_Mix15_3 | QC_LowIS_Mix16_3 | QC_LowIS_Mix17_3 | QC_LowIS_Mix18_3 | QC_LowIS_Mix19_3 | QC_LowIS_Mix20_3
LowIS_Mix18_2 | HighIS_Mix14_2 | LowIS_Mix06_1 | 1 | 6 | QC_LowIS_Mix18_2 | QC_HighIS_Mix20_1 | QC_LowIS_Mix16_3 |  | 2 | 30 | D
LowIS_Mix18_3 | LowIS_Mix14_1 | LowIS_Mix05_3 | 1 | 29 | QC_LowIS_Mix18_3 | QC_HighIS_Mix11_1 | QC_HighIS_Mix15_1 |  | 2 | 65 | E
LowIS_Mix19_1 | HighIS_Mix08_3 | LowIS_Mix01_2 | 1 | 13 | QC_LowIS_Mix19_1 | QC_LowIS_Mix09_1 | QC_HighIS_Mix16_3 |  | 2 | 90 | F | QC_HighIS_Mix11_1 | QC_HighIS_Mix12_1 | QC_HighIS_Mix13_1 | QC_HighIS_Mix14_1 | QC_HighIS_Mix15_1 | QC_HighIS_Mix16_1 | QC_HighIS_Mix17_1 | QC_HighIS_Mix18_1 | QC_HighIS_Mix19_1 | QC_HighIS_Mix20_1
LowIS_Mix19_2 | HighIS_Mix13_1 | HighIS_Mix09_3 | 1 | 93 | QC_LowIS_Mix19_2 | QC_LowIS_Mix02_3 | QC_HighIS_Mix20_3 |  | 2 | 94 | G | QC_HighIS_Mix11_2 | QC_HighIS_Mix12_2 | QC_HighIS_Mix13_2 | QC_HighIS_Mix14_2 | QC_HighIS_Mix15_2 | QC_HighIS_Mix16_2 | QC_HighIS_Mix17_2 | QC_HighIS_Mix18_2 | QC_HighIS_Mix19_2 | QC_HighIS_Mix20_2
LowIS_Mix19_3 | HighIS_Mix16_2 | LowIS_Mix17_1 | 2 | 7 | QC_LowIS_Mix19_3 | QC_HighIS_Mix07_3 | QC_LowIS_Mix09_3 |  | 1 | 33 | H | QC_HighIS_Mix11_3 | QC_HighIS_Mix12_3 | QC_HighIS_Mix13_3 | QC_HighIS_Mix14_3 | QC_HighIS_Mix15_3 | QC_HighIS_Mix16_3 | QC_HighIS_Mix17_3 | QC_HighIS_Mix18_3 | QC_HighIS_Mix19_3 | QC_HighIS_Mix20_3
LowIS_Mix20_1 | HighIS_Mix08_2 | HighIS_Mix16_1 | 2 | 66 | QC_LowIS_Mix20_1 | QC_LowIS_Mix18_2 | QC_LowIS_Mix15_1 |  | 2 | 5
LowIS_Mix20_2 | HighIS_Mix16_3 | LowIS_Mix11_3 | 2 | 25 | QC_LowIS_Mix20_2 | QC_HighIS_Mix13_1 | QC_HighIS_Mix14_2 |  | 2 | 76
LowIS_Mix20_3 | HighIS_Mix19_2 | LowIS_Mix10_2 | 1 | 22 | QC_LowIS_Mix20_3 | QC_LowIS_Mix05_1 | QC_HighIS_Mix18_2 |  | 2 | 80
HighIS_Mix01_1 | LowIS_Mix12_1 | HighIS_Mix05_2 | 1 | 77 | QC_HighIS_Mix01_1 | QC_HighIS_Mix10_1 | QC_LowIS_Mix13_2 |  | 2 | 15
HighIS_Mix01_2 | LowIS_Mix18_2 | HighIS_Mix14_2 | 2 | 76 | QC_HighIS_Mix01_2 | QC_LowIS_Mix19_2 | QC_HighIS_Mix17_1 |  | 2 | 67
HighIS_Mix01_3 | HighIS_Mix14_3 | LowIS_Mix14_1 | 2 | 4 | QC_HighIS_Mix01_3 | QC_HighIS_Mix12_2 | QC_HighIS_Mix09_2 |  | 1 | 81
HighIS_Mix02_1 | HighIS_Mix18_1 | HighIS_Mix08_3 | 1 | 92 | QC_HighIS_Mix02_1 | QC_LowIS_Mix10_2 | QC_LowIS_Mix11_3 |  | 2 | 25
HighIS_Mix02_2 | HighIS_Mix01_1 | HighIS_Mix13_1 | 2 | 63 | QC_HighIS_Mix02_2 | QC_HighIS_Mix16_1 | QC_LowIS_Mix03_2 |  | 1 | 15
HighIS_Mix02_3 | LowIS_Mix01_1 | HighIS_Mix16_2 | 2 | 78 | QC_HighIS_Mix02_3 | QC_HighIS_Mix06_2 | QC_LowIS_Mix16_2 |  | 2 | 18
HighIS_Mix03_1 | LowIS_Mix16_2 | HighIS_Mix08_2 | 1 | 80 | QC_HighIS_Mix03_1 | QC_LowIS_Mix19_3 | QC_HighIS_Mix12_3 |  | 2 | 86
HighIS_Mix03_2 | HighIS_Mix03_2 | HighIS_Mix16_3 | 2 | 90 | QC_HighIS_Mix03_2 | QC_LowIS_Mix02_1 | QC_HighIS_Mix20_1 |  | 2 | 70
HighIS_Mix03_3 | LowIS_Mix10_1 | HighIS_Mix19_2 | 2 | 81 | QC_HighIS_Mix03_3 | QC_HighIS_Mix05_1 | QC_HighIS_Mix11_1 |  | 2 | 61
HighIS_Mix04_1 | HighIS_Mix07_1 | LowIS_Mix12_1 | 2 | 2 | QC_HighIS_Mix04_1 | QC_LowIS_Mix12_3 | QC_LowIS_Mix09_1 |  | 1 | 9
HighIS_Mix04_2 | HighIS_Mix07_2 | LowIS_Mix18_2 | 2 | 20 | QC_HighIS_Mix04_2 | QC_LowIS_Mix07_3 | QC_LowIS_Mix02_3 |  | 1 | 26
HighIS_Mix04_3 | LowIS_Mix05_2 | HighIS_Mix14_3 | 2 | 88 | QC_HighIS_Mix04_3 | QC_HighIS_Mix03_3 | QC_HighIS_Mix07_3 |  | 1 | 91
HighIS_Mix05_1 | HighIS_Mix06_1 | HighIS_Mix18_1 | 2 | 68 | QC_HighIS_Mix05_1 | QC_LowIS_Mix16_1 | QC_LowIS_Mix18_2 |  | 2 | 20
HighIS_Mix05_2 | LowIS_Mix11_1 | HighIS_Mix01_1 | 1 | 61 | QC_HighIS_Mix05_2 | QC_HighIS_Mix18_1 | QC_HighIS_Mix13_1 |  | 2 | 63
HighIS_Mix05_3 | HighIS_Mix12_1 | LowIS_Mix01_1 | 1 | 1 | QC_HighIS_Mix05_3 | QC_LowIS_Mix20_1 | QC_LowIS_Mix05_1 |  | 1 | 5
HighIS_Mix06_1 | HighIS_Mix11_1 | LowIS_Mix16_2 | 2 | 18 | QC_HighIS_Mix06_1 | QC_HighIS_Mix05_3 | QC_HighIS_Mix10_1 |  | 1 | 70
HighIS_Mix06_2 | LowIS_Mix03_1 | HighIS_Mix03_2 | 1 | 75 | QC_HighIS_Mix06_2 | QC_HighIS_Mix16_2 | QC_LowIS_Mix19_2 |  | 2 | 21
HighIS_Mix06_3 | HighIS_Mix19_1 | LowIS_Mix10_1 | 1 | 10 | QC_HighIS_Mix06_3 | QC_LowIS_Mix06_3 | QC_HighIS_Mix12_2 |  | 2 | 74
HighIS_Mix07_1 | LowIS_Mix04_3 | HighIS_Mix07_1 | 1 | 67 | QC_HighIS_Mix07_1 | QC_LowIS_Mix04_3 | QC_LowIS_Mix10_2 |  | 1 | 22
HighIS_Mix07_2 | HighIS_Mix06_3 | HighIS_Mix07_2 | 1 | 79 | QC_HighIS_Mix07_2 | QC_HighIS_Mix01_3 | QC_HighIS_Mix16_1 |  | 2 | 66
HighIS_Mix07_3 | HighIS_Mix15_3 | LowIS_Mix05_2 | 1 | 17 | QC_HighIS_Mix07_3 | QC_LowIS_Mix14_3 | QC_HighIS_Mix06_2 |  | 1 | 78
HighIS_Mix08_1 | LowIS_Mix03_2 | HighIS_Mix06_1 | 1 | 66 | QC_HighIS_Mix08_1 | QC_HighIS_Mix08_2 | QC_LowIS_Mix19_3 |  | 2 | 33
HighIS_Mix08_2 | LowIS_Mix18_1 | LowIS_Mix11_1 | 2 | 1 | QC_HighIS_Mix08_2 | QC_HighIS_Mix02_2 | QC_LowIS_Mix02_1 |  | 1 | 2
HighIS_Mix08_3 | HighIS_Mix17_1 | HighIS_Mix12_1 | 2 | 62 | QC_HighIS_Mix08_3 | QC_HighIS_Mix15_3 | QC_HighIS_Mix05_1 |  | 1 | 65
HighIS_Mix09_1 | LowIS_Mix09_2 | HighIS_Mix11_1 | 2 | 61 | QC_HighIS_Mix09_1 | QC_HighIS_Mix07_2 | QC_LowIS_Mix12_3 |  | 2 | 26
HighIS_Mix09_2 | LowIS_Mix08_1 | LowIS_Mix03_1 | 1 | 3 | QC_HighIS_Mix09_2 | QC_LowIS_Mix05_3 | QC_LowIS_Mix07_3 |  | 1 | 31
HighIS_Mix09_3 | HighIS_Mix10_1 | HighIS_Mix19_1 | 2 | 69 | QC_HighIS_Mix09_3 | QC_LowIS_Mix17_1 | QC_HighIS_Mix03_3 |  | 1 | 87
HighIS_Mix10_1 | LowIS_Mix17_2 | LowIS_Mix04_3 | 1 | 28 | QC_HighIS_Mix10_1 | QC_HighIS_Mix06_1 | QC_LowIS_Mix16_1 |  | 2 | 6
HighIS_Mix10_2 | LowIS_Mix01_3 | HighIS_Mix06_3 | 1 | 90 | QC_HighIS_Mix10_2 | QC_HighIS_Mix03_2 | QC_HighIS_Mix18_1 |  | 2 | 68
HighIS_Mix10_3 | LowIS_Mix04_2 | HighIS_Mix15_3 | 2 | 89 | QC_HighIS_Mix10_3 | QC_HighIS_Mix19_1 | QC_LowIS_Mix20_1 |  | 2 | 10
HighIS_Mix11_1 | LowIS_Mix15_3 | LowIS_Mix03_2 | 1 | 15 | QC_HighIS_Mix11_1 | QC_LowIS_Mix19_1 | QC_HighIS_Mix05_3 |  | 1 | 89
HighIS_Mix11_2 | HighIS_Mix11_3 | LowIS_Mix18_1 | 2 | 8 | QC_HighIS_Mix11_2 | QC_LowIS_Mix12_2 | QC_HighIS_Mix16_2 |  | 2 | 78
HighIS_Mix11_3 | LowIS_Mix17_3 | HighIS_Mix17_1 | 2 | 67 | QC_HighIS_Mix11_3 | QC_HighIS_Mix02_3 | QC_LowIS_Mix06_3 |  | 1 | 30
HighIS_Mix12_1 | LowIS_Mix20_3 | LowIS_Mix09_2 | 1 | 21 | QC_HighIS_Mix12_1 | QC_HighIS_Mix14_1 | QC_LowIS_Mix04_3 |  | 1 | 28
HighIS_Mix12_2 | HighIS_Mix07_3 | LowIS_Mix08_1 | 1 | 8 | QC_HighIS_Mix12_2 | QC_LowIS_Mix05_2 | QC_HighIS_Mix01_3 |  | 1 | 85
HighIS_Mix12_3 | LowIS_Mix08_3 | HighIS_Mix10_1 | 1 | 70 | QC_HighIS_Mix12_3 | QC_HighIS_Mix01_2 | QC_LowIS_Mix14_3 |  | 2 | 28
HighIS_Mix13_1 | LowIS_Mix18_3 | LowIS_Mix17_2 | 2 | 19 | QC_HighIS_Mix13_1 | QC_HighIS_Mix07_1 | QC_HighIS_Mix08_2 |  | 1 | 80
HighIS_Mix13_2 | HighIS_Mix20_2 | LowIS_Mix01_3 | 1 | 25 | QC_HighIS_Mix13_2 | QC_HighIS_Mix09_3 | QC_HighIS_Mix02_2 |  | 1 | 74
HighIS_Mix13_3 | LowIS_Mix13_1 | LowIS_Mix04_2 | 1 | 16 | QC_HighIS_Mix13_3 | QC_LowIS_Mix20_2 | QC_HighIS_Mix15_3 |  | 2 | 89
HighIS_Mix14_1 | LowIS_Mix09_3 | LowIS_Mix15_3 | 2 | 29 | QC_HighIS_Mix14_1 | QC_LowIS_Mix11_2 | QC_HighIS_Mix07_2 |  | 1 | 79
HighIS_Mix14_2 | LowIS_Mix19_1 | HighIS_Mix11_3 | 2 | 85 | QC_HighIS_Mix14_2 | QC_HighIS_Mix12_1 | QC_LowIS_Mix05_3 |  | 1 | 29
HighIS_Mix14_3 | HighIS_Mix01_2 | LowIS_Mix17_3 | 2 | 31 | QC_HighIS_Mix14_3 | QC_LowIS_Mix12_1 | QC_LowIS_Mix17_1 |  | 2 | 7
HighIS_Mix15_1 | HighIS_Mix05_1 | LowIS_Mix20_3 | 2 | 34 | QC_HighIS_Mix15_1 | QC_LowIS_Mix01_3 | QC_HighIS_Mix06_1 |  | 1 | 66
HighIS_Mix15_2 | LowIS_Mix05_1 | HighIS_Mix07_3 | 1 | 91 | QC_HighIS_Mix15_2 | QC_HighIS_Mix15_2 | QC_HighIS_Mix03_2 |  | 1 | 75
HighIS_Mix15_3 | HighIS_Mix18_2 | LowIS_Mix08_3 | 1 | 32 | QC_HighIS_Mix15_3 | QC_LowIS_Mix06_1 | QC_HighIS_Mix19_1 |  | 2 | 69
HighIS_Mix16_1 | HighIS_Mix17_2 | LowIS_Mix18_3 | 2 | 32 | QC_HighIS_Mix16_1 | QC_LowIS_Mix03_3 | QC_LowIS_Mix19_1 |  | 2 | 9
HighIS_Mix16_2 | LowIS_Mix03_3 | HighIS_Mix20_2 | 2 | 82 | QC_HighIS_Mix16_2 | QC_HighIS_Mix11_3 | QC_LowIS_Mix12_2 |  | 2 | 14
HighIS_Mix16_3 | LowIS_Mix02_1 | LowIS_Mix13_1 | 2 | 3 | QC_HighIS_Mix16_3 | QC_HighIS_Mix18_3 | QC_HighIS_Mix02_3 |  | 1 | 86
HighIS_Mix17_1 | LowIS_Mix20_1 | LowIS_Mix09_3 | 1 | 33 | QC_HighIS_Mix17_1 | QC_LowIS_Mix15_2 | QC_HighIS_Mix14_1 |  | 2 | 64
HighIS_Mix17_2 | HighIS_Mix10_3 | LowIS_Mix19_1 | 2 | 9 | QC_HighIS_Mix17_2 | QC_LowIS_Mix06_2 | QC_LowIS_Mix05_2 |  | 1 | 17
HighIS_Mix17_3 | HighIS_Mix04_2 | HighIS_Mix01_2 | 1 | 73 | QC_HighIS_Mix17_3 | QC_LowIS_Mix10_3 | QC_HighIS_Mix01_2 |  | 1 | 73
HighIS_Mix18_1 | LowIS_Mix06_2 | HighIS_Mix05_1 | 1 | 65 | QC_HighIS_Mix18_1 | QC_LowIS_Mix07_2 | QC_HighIS_Mix07_1 |  | 1 | 67
HighIS_Mix18_2 | LowIS_Mix15_2 | LowIS_Mix05_1 | 1 | 5 | QC_HighIS_Mix18_2 | QC_HighIS_Mix19_2 | QC_HighIS_Mix09_3 |  | 1 | 93
HighIS_Mix18_3 | LowIS_Mix15_1 | HighIS_Mix18_2 | 2 | 80 | QC_HighIS_Mix18_3 | QC_LowIS_Mix01_1 | QC_LowIS_Mix20_2 |  | 2 | 22
HighIS_Mix19_1 | HighIS_Mix15_1 | HighIS_Mix17_2 | 2 | 79 | QC_HighIS_Mix19_1 | QC_HighIS_Mix02_1 | QC_LowIS_Mix11_2 |  | 2 | 13
HighIS_Mix19_2 | HighIS_Mix04_3 | LowIS_Mix03_3 | 1 | 27 | QC_HighIS_Mix19_2 | QC_HighIS_Mix01_1 | QC_HighIS_Mix12_1 |  | 2 | 62
HighIS_Mix19_3 | HighIS_Mix20_3 | LowIS_Mix02_1 | 1 | 2 | QC_HighIS_Mix19_3 | QC_HighIS_Mix06_3 | QC_LowIS_Mix12_1 |  | 2 | 2
HighIS_Mix20_1 | LowIS_Mix16_1 | LowIS_Mix20_1 | 2 | 10 | QC_HighIS_Mix20_1 | QC_HighIS_Mix05_2 | QC_LowIS_Mix01_3 |  | 1 | 25
HighIS_Mix20_2 | HighIS_Mix05_3 | HighIS_Mix10_3 | 1 | 94 | QC_HighIS_Mix20_2 | QC_LowIS_Mix20_3 | QC_HighIS_Mix15_2 |  | 2 | 77
HighIS_Mix20_3 | HighIS_Mix13_2 | HighIS_Mix04_2 | 1 | 76 | QC_HighIS_Mix20_3 | QC_LowIS_Mix17_3 | QC_LowIS_Mix06_1 |  | 1 | 6
|  | LowIS_Mix06_2 | 1 | 18 |  |  | QC_LowIS_Mix03_3 |  | 1 | 27
|  | LowIS_Mix15_2 | 2 | 17 |  |  | QC_HighIS_Mix11_3 |  | 2 | 85
|  | LowIS_Mix15_1 | 2 | 5 |  |  | QC_HighIS_Mix18_3 |  | 2 | 92
|  | HighIS_Mix15_1 | 2 | 65 |  |  | QC_LowIS_Mix15_2 |  | 2 | 17
|  | HighIS_Mix04_3 | 1 | 88 |  |  | QC_LowIS_Mix06_2 |  | 1 | 18
|  | HighIS_Mix20_3 | 2 | 94 |  |  | QC_LowIS_Mix10_3 |  | 1 | 34
|  | LowIS_Mix16_1 | 2 | 6 |  |  | QC_LowIS_Mix07_2 |  | 1 | 19
|  | HighIS_Mix05_3 | 1 | 89 |  |  | QC_HighIS_Mix19_2 |  | 2 | 81
|  | HighIS_Mix13_2 | 2 | 75 |  |  | QC_LowIS_Mix01_1 |  | 1 | 1
|  | HighIS_Mix01_CE20 | 1 | 73 |  |  | QC_HighIS_Mix02_1 |  | 1 | 62
|  | HighIS_Mix01_CE30 | 1 | 73 |  |  | QC_HighIS_Mix01_1 |  | 1 | 61
|  | HighIS_Mix02_CE20 | 1 | 74 |  |  | QC_HighIS_Mix06_3 |  | 1 | 90
|  | HighIS_Mix02_CE30 | 1 | 74 |  |  | QC_HighIS_Mix05_2 |  | 1 | 77
|  | HighIS_Mix03_CE20 | 1 | 75 |  |  | QC_LowIS_Mix20_3 |  | 2 | 34
|  | HighIS_Mix03_CE30 | 1 | 75 |  |  | QC_LowIS_Mix17_3 |  | 2 | 31
|  | HighIS_Mix04_CE20 | 1 | 76 |  |  | QC_HighIS_Mix01_CE20 |  | 1 | 73
|  | HighIS_Mix04_CE30 | 1 | 76 |  |  | QC_HighIS_Mix01_CE30 |  | 1 | 73
|  | HighIS_Mix05_CE20 | 1 | 77 |  |  | QC_HighIS_Mix02_CE20 |  | 1 | 74
|  | HighIS_Mix05_CE30 | 1 | 77 |  |  | QC_HighIS_Mix02_CE30 |  | 1 | 74
|  | HighIS_Mix06_CE20 | 1 | 78 |  |  | QC_HighIS_Mix03_CE20 |  | 1 | 75
|  | HighIS_Mix06_CE30 | 1 | 78 |  |  | QC_HighIS_Mix03_CE30 |  | 1 | 75
|  | HighIS_Mix07_CE20 | 1 | 79 |  |  | QC_HighIS_Mix04_CE20 |  | 1 | 76
|  | HighIS_Mix07_CE30 | 1 | 79 |  |  | QC_HighIS_Mix04_CE30 |  | 1 | 76
|  | HighIS_Mix08_CE20 | 1 | 80 |  |  | QC_HighIS_Mix05_CE20 |  | 1 | 77
|  | HighIS_Mix08_CE30 | 1 | 80 |  |  | QC_HighIS_Mix05_CE30 |  | 1 | 77
|  | HighIS_Mix09_CE20 | 1 | 81 |  |  | QC_HighIS_Mix06_CE20 |  | 1 | 78
|  | HighIS_Mix09_CE30 | 1 | 81 |  |  | QC_HighIS_Mix06_CE30 |  | 1 | 78
|  | HighIS_Mix10_CE20 | 1 | 82 |  |  | QC_HighIS_Mix07_CE20 |  | 1 | 79
|  | HighIS_Mix10_CE30 | 1 | 82 |  |  | QC_HighIS_Mix07_CE30 |  | 1 | 79
|  | HighIS_Mix11_CE20 | 2 | 73 |  |  | QC_HighIS_Mix08_CE20 |  | 1 | 80
|  | HighIS_Mix11_CE30 | 2 | 73 |  |  | QC_HighIS_Mix08_CE30 |  | 1 | 80
|  | HighIS_Mix12_CE20 | 2 | 74 |  |  | QC_HighIS_Mix09_CE20 |  | 1 | 81
|  | HighIS_Mix12_CE30 | 2 | 74 |  |  | QC_HighIS_Mix09_CE30 |  | 1 | 81
|  | HighIS_Mix13_CE20 | 2 | 75 |  |  | QC_HighIS_Mix10_CE20 |  | 1 | 82
|  | HighIS_Mix13_CE30 | 2 | 75 |  |  | QC_HighIS_Mix10_CE30 |  | 1 | 82
|  | HighIS_Mix14_CE20 | 2 | 76 |  |  | QC_HighIS_Mix11_CE20 |  | 2 | 73
|  | HighIS_Mix14_CE30 | 2 | 76 |  |  | QC_HighIS_Mix11_CE30 |  | 2 | 73
|  | HighIS_Mix15_CE20 | 2 | 77 |  |  | QC_HighIS_Mix12_CE20 |  | 2 | 74
|  | HighIS_Mix15_CE30 | 2 | 77 |  |  | QC_HighIS_Mix12_CE30 |  | 2 | 74
|  | HighIS_Mix16_CE20 | 2 | 78 |  |  | QC_HighIS_Mix13_CE20 |  | 2 | 75
|  | HighIS_Mix16_CE30 | 2 | 78 |  |  | QC_HighIS_Mix13_CE30 |  | 2 | 75
|  | HighIS_Mix17_CE20 | 2 | 79 |  |  | QC_HighIS_Mix14_CE20 |  | 2 | 76
|  | HighIS_Mix17_CE30 | 2 | 79 |  |  | QC_HighIS_Mix14_CE30 |  | 2 | 76
|  | HighIS_Mix18_CE20 | 2 | 80 |  |  | QC_HighIS_Mix15_CE20 |  | 2 | 77
|  | HighIS_Mix18_CE30 | 2 | 80 |  |  | QC_HighIS_Mix15_CE30 |  | 2 | 77
|  | HighIS_Mix19_CE20 | 2 | 81 |  |  | QC_HighIS_Mix16_CE20 |  | 2 | 78
|  | HighIS_Mix19_CE30 | 2 | 81 |  |  | QC_HighIS_Mix16_CE30 |  | 2 | 78
|  | HighIS_Mix20_CE20 | 2 | 82 |  |  | QC_HighIS_Mix17_CE20 |  | 2 | 79
|  | HighIS_Mix20_CE30 | 2 | 82 |  |  | QC_HighIS_Mix17_CE30 |  | 2 | 79
[7 empty] | QC_HighIS_Mix18_CE20 |  | 2 | 80
[7 empty] | QC_HighIS_Mix18_CE30 |  | 2 | 80
[7 empty] | QC_HighIS_Mix19_CE20 |  | 2 | 81
[7 empty] | QC_HighIS_Mix19_CE30 |  | 2 | 81
[7 empty] | QC_HighIS_Mix20_CE20 |  | 2 | 82
[7 empty] | QC_HighIS_Mix20_CE30 |  | 2 | 82

## Sample Preparation
Sample Preparation - All Mixes Experiment
LowIS - Final Concentration: 5ppm
For each mix: 3LowIS and 3 for QC_LowIS
Each sample: 200uL of solution
Total volume needed: 6 x 200uL = 1200uL (prepare 4mL)
From 1mg/mL stock solution, pippete 20uL
C1V1 = C2V2
1000ppm x 20uL = C2 x 4000uL
C2 = 5ppm
HighIS - Final concetration: 50ppm
For each mix: 3HighIS and 3 for QC_HighIS
Each sample: 200uL of solution
Total volume needed: 6 x 200uL = 1200uL (prepare 4mL)
From 1mg/mL stock solution, pippete 200uL
C1V1 = C2V2
1000ppm x 200uL = C2 x 4000uL
C2 = 50ppm
After pipetting each standard, complete with Water/ACN until 2mL. Attention: this volume will change depending on the number of metabolites inside each mix
DAY 1 - Monday, January 13rd, 2020
1 | Take stock solutions from -80°C and let defrost in the fridge (whole morning - DMSO takes more time to defrost)
2 | Prepare 2 bottles of 2L from each of the following:
Mobile Phase A1: ACN + 0,1% Formic Acid
Mobile Phase B1: Water MilliQ + 0,1% Formic Acid
Mobile Phase A2: ACN 100%
Mobile Phase B2: Water MilliQ
Auto sampler needle wash solvents (ACN: H20:IPA
* Each bottle will be used for one polarity. So, after running Water/ACN batch in positive mode, bottles will be changed to the 2nd one to run negative mode. In this way, the same bottle is used for both Water/ACN and Serum batches, avoiding RT shifts due to the variation of the mobile phase
3 | Change column for a new one (pre-column is already new) and wash for 1 hour at 0.2mL.min-1 with initial method gradient
4 | Clean the ion source
5 | Perform instrument optimization in positive mode
6 | Prepare worklist files
7 | Take stock solutions from the fridge and let them for 30min at room temperature
8 | In the meantime, purge and condition the LC system and start the pumps under low flow (0.1mL.min-1)
9 | Prepare LowIS and HighIS solutions according to the table above
10 | Fill the plates according to the map and centrifuge for 30min at 4°C (this is the temperature of the autosampler at the equipment. Adding this step avoid that unexpected precipitation due to low temperature occur in the autosampler with no centrifugation, blocking the needle
11 | In the meantime, refill calibrant solutions and and calibrate instrument in positive mode
12 | Take the plates to the autosampler and start the run. Estimated Time: 25h (169 injections x ~9min)
13 | Prepare serum pool for next day (123 samples x 50uL = 6.15mL). Take 10mL
DAY 2 - Tuesday, January 14th, 2020
1 | Prepare Serum batch samples (untargeted protocol) and store at -80°C
2 | When POS is finished, change mobile phase, purge the system, optimize, calibrate and start Water/ACN batch in NEG mode
DAY 3 - Wednesday, January 15th, 2020
1 | A bit before Water/ACN batch analysis in NEG mode finishes, re-suspend POS mode serum plates (as usual, POS and NEG are splitted in the end of the sample preparation protocol)
2 | When Water/ACN batch analysis in NEG mdoe finishes, change mobile phase, purge, clean the source, calibrate for POS and start serum batch in POS mode
DAY 4 - Thursday, January 16th, 2020
1 | A bit before serum batch analysis in POS mode finishes, re-suspend NEG mode serum plates
2 | When serum batch analysis in POS mode finishes, change mobile phase, purge, calibrate for NEG and start serum batch in NEG mode
DAY 5 - Friday, January 17th, 2020
1 | Copy wiff files to massspec folder
2 | Clean the source and column with no aditives